## Reference Materials
MEDALS D4.6 – Reference Materials and European Material Grades
Reference compositions and published grade specifications used to interpret the measured LIBS sub-datasets. BAS/BCS values are reference compositions; European grade values are published specifications/typical compositions and are not sample-specific measurements.
Legend: * = approximate BAS value rather than certified value. “Balance” denotes the remaining matrix element. For the phosphor-bronze series, values with “<” retain the reporting qualifier. European material-grade compositions are reference specifications/typical values and must not be treated as measured ground truth for the individual MEDALS samples.
A. BAS Ferritic Stainless-Steel Reference Materials (469–473)
Sample | Reference material | Matrix | C | Si | Mn | P | S | Cr | Ni | Mo | Co | Cu | V | Fe | Composition status | Notes | Source
469 | Ferritic stainless steel | Fe-based | 0.279 | 0.421 | 0.598 | 0.015 | 0.02 | 11.93 | 0.246 | — | 0.01* | 0.02* | 0.02* | Balance | Reference composition | * values approximate | S1
470 | Ferritic stainless steel | Fe-based | 0.153 | 0.335 | 0.235 | 0.024 | 0.035 | 17.68 | 0.369 | — | 0.02* | 0.02* | 0.02* | Balance | Reference composition | * values approximate | S1
471 | Ferritic stainless steel | Fe-based | 0.095 | 0.326 | 0.417 | 0.018 | 0.023 | 23.85 | 0.96 | — | 0.02* | 0.02* | 0.03* | Balance | Reference composition | * values approximate | S1
472 | Ferritic stainless steel | Fe-based | 0.227 | 1.05 | 1.02 | 0.032 | 0.029 | 15.82 | 1.95 | 0.661 | 0.02* | 0.02* | 0.02* | Balance | Reference composition | * values approximate | S1
473 | Ferritic stainless steel | Fe-based | 0.172 | 0.604 | 0.494 | 0.019 | 0.03 | 9.06 | 0.060* | 0.95 | 0.01* | 0.03* | 0.02* | Balance | Reference composition | * values approximate | S1
B. BAS/BCS Phosphor-Bronze Reference Materials (551–556)
Sample | Reference material | Matrix | Sn | P | Pb | Ni | Zn | Cu | Si | Fe | Al | Composition status | Source
551 | Phosphor bronze | Cu-based | 8.92 | 1.01 | 0.79 | 0.76 | 0.74 | 87.4 | 0.018 | 0.2 | 0.052 | Reference composition | S2
552 | Phosphor bronze | Cu-based | 9.78 | 0.77 | 0.63 | 0.56 | 0.35 | 87.7 | 0.019 | 0.1 | 0.023 | Reference composition | S2
553 | Phosphor bronze | Cu-based | 10.8 | 0.68 | 0.47 | 0.44 | 0.49 | 87 | 0.022 | 0.056 | 0.017 | Reference composition | S2
554 | Phosphor bronze | Cu-based | 11.3 | 0.41 | 0.34 | 0.22 | 0.22 | 87.4 | 0.038 | 0.022 | 0.005 | Reference composition | S2
555 | Phosphor bronze | Cu-based | 12.1 | 0.18 | 0.24 | 0.11 | 0.16 | 87.1 | 0.036 | 0.01 | <0.005 | Reference composition | S2
556 | Phosphor bronze | Cu-based | 13.2 | 0.1 | 0.16 | 0.014 | 0.09 | 86.4 | <0.005 | 0.004 | <0.005 | Reference composition | S2
C. European Material Grades Represented in the Dataset
Material no. | Common designation | Material family | C | Si | Mn | P | S | N | Cr | Ni | Mo | Cu | Nb | Al | Co | Fe | Source / interpretation
1.4301 | 304 / X5CrNi18-10 | Austenitic stainless steel | ≤0.07 | ≤1.00 | ≤2.00 | ≤0.045 | ≤0.015 | ≤0.11 | 17.5–19.5 | 8.0–10.5 | — | — | — | — | — | Balance | S3 – BS EN 10088-2 composition limits
1.4404 | 316L / X2CrNiMo17-12-2 | Austenitic stainless steel | ≤0.030 | ≤1.00 | ≤2.00 | ≤0.045 | ≤0.015 | ≤0.10 | 16.5–18.5 | 10.0–13.0 | 2.00–3.00 | — | — | — | — | Balance | S3 – BS EN 10088-2 composition limits
1.4462 | 2205 / X2CrNiMoN22-5-3 | Duplex stainless steel | ≤0.030 | ≤1.00 | ≤2.00 | ≤0.035 | ≤0.015 | 0.10–0.22 | 21.0–23.0 | 4.5–6.5 | 2.50–3.50 | — | — | — | — | Balance | S4 – BS EN 10088-3 composition limits
1.4539 | 904L / X1NiCrMoCu25-20-5 | Highly alloyed austenitic stainless steel | ≤0.020 | ≤0.70 | ≤2.00 | ≤0.030 | ≤0.010 | ≤0.15 | 19.0–21.0 | 24.0–26.0 | 4.0–5.0 | 1.20–2.00 | — | — | — | Balance | S4 – BS EN 10088-3 composition limits
1.3974 | X2CrNiMnMoNNb23-17-6-3 | Non-magnetic special steel | ≤0.03 | ≤1.00 | 4.50–6.50 | ≤0.025 | ≤0.010 | 0.30–0.50 | 21.0–24.5 | 15.5–18.0 | 2.80–3.40 | — | 0.10–0.30 | — | — | Balance | S5 – Walzwerke Einsal material data
1.4860 | NiCr 30 20 | Fe–Ni–Cr heating conductor alloy | — | — | — | — | — | — | 20 (avg.) | 30 (avg.) | — | — | — | — | — | Balance | S6 – DIN 17470 / BGH average composition; not sample-specific
2.4605 | Alloy 59 / UNS N06059 / NiCr23Mo16Al | Nickel-based corrosion-resistant alloy | ≤0.01 | ≤0.10 | ≤0.5 | ≤0.015 | ≤0.01 | — | 22.0–24.0 | Balance | 15.0–16.5 | ≤0.5 | — | 0.10–0.40 | ≤0.3 | ≤1.5 | S7 – VDM Alloy 59 specification

## Feb26
Sub-dataset – Medals_Samples_Feb26
Measured LIBS spectra from 16 sample IDs. The folder name contains “Feb26”, while the acquisition date encoded in filenames is 11 March 2026.
Dataset overview
Property | Specification
Repository location | data/LSA/Medals_Samples_Feb26
Sample folders | 16
Measurement files per sample | 5
Spectra per measurement file | 40
Spectra per sample | 200
Total spectral files | 80 CSVs
Total spectra | 3200
Wavelength channels per spectrum | 2048
Total intensity values | 6553600
Supporting files | 1 composition JSON, 1 settings TXT, 3 JPEG images
Total directory size | Approximately 31.85 MiB
Acquisition date in filenames | 11 March 2026
Filename time span | 16:20:55–16:47:45; timezone unspecified
Organisation | Per-sample folders; 5 CSV measurement files per sample
Sample inventory
Sample ID | Material information | Files | Spectra per file | Total spectra | Ground-truth composition | Notes
469 | Ferritic stainless-steel reference material | 5 | 40 | 200 | Available | Reference composition in metadata
470 | Ferritic stainless-steel reference material | 5 | 40 | 200 | Available | Reference composition in metadata
471 | Ferritic stainless-steel reference material | 5 | 40 | 200 | Available | Reference composition in metadata
472 | Ferritic stainless-steel reference material | 5 | 40 | 200 | Available | Reference composition in metadata
473 | Ferritic stainless-steel reference material | 5 | 40 | 200 | Available | Reference composition in metadata
551 | Phosphor-bronze reference material | 5 | 40 | 200 | Available | Reference composition in metadata
552 | Phosphor-bronze reference material | 5 | 40 | 200 | Available | Reference composition in metadata
553 | Phosphor-bronze reference material | 5 | 40 | 200 | Available | Reference composition in metadata
554 | Phosphor-bronze reference material | 5 | 40 | 200 | Available | Reference composition in metadata
555 | Phosphor-bronze reference material | 5 | 40 | 200 | Available | Reference composition in metadata
556 | Phosphor-bronze reference material | 5 | 40 | 200 | Available | Reference composition in metadata
Al_ENAS-42000-1 | Aluminium-labelled sample | 5 | 40 | 200 | Not provided in folder JSON
M2-7_A3 | Sample label only | 5 | 40 | 200 | Not provided
SS405 | Sample label only | 5 | 40 | 200 | Not provided
SS465-1 | Sample label only | 5 | 40 | 200 | Not provided
SS467 | Sample label only | 5 | 40 | 200 | Not provided
CSV format and spectral axis
Property | Specification
Separator | Semicolon (;)
Header rows | 1
Data rows per CSV | 40
Columns per CSV | 2050
Metadata columns | MeasCtr, TimeStamp
MeasCtr | Consistently 1–40 in every file
TimeStamp | Integer acquisition timestamp; unit and origin not documented
Spectral intensity columns | 2048
Wavelength coverage | 341.952–896.916 nm
Wavelength ordering | Strictly increasing; identical across all 80 files
Channel spacing | Nonuniform, approximately 0.242–0.294 nm
Mean channel spacing | Approximately 0.2711 nm
Intensity representation | Integer values; physical intensity unit not documented
Observed intensity range | 496–65,535
Optical spectral resolution | Not specified; must not be inferred from channel spacing
Acquisition hardware and recorded settings
Property | Specification
Spectrometer manufacturer | Avantes
Spectrometer designation | AS7010
Channels | 2048
Instrument identifier | 7518629SP
Exposure time | 0.1 ms (100 µs)
Laser | RevAmp
Pulse divider | 2
Stated pulse operation | 25 Hz, single pulse
Current setting | 630 A
Not specified | Laser wavelength, pulse energy, pulse duration, spot size, collection geometry, trigger delay
Timing caution | Most adjacent TimeStamp values differ by 20 units. Timestamp unit/synchronisation is not documented, so shot timing should not be inferred without confirmation.
Quality checks and analysis implications
Check / item | Result / implication
Consistent structure | All files have 40 rows and the same 2,048-channel wavelength grid.
Missing/malformed intensity data | None found; all intensities parsed as integers.
Exact duplicate CSV files | None; all 80 CSVs have distinct content hashes.
Intensity values equal to 65,535 | 1,641 values
Spectra containing ≥1 value of 65,535 | 492 of 3,200 spectra (15.4%)
Saturation interpretation | Repeated maximum unsigned 16-bit value strongly suggests clipping at some wavelengths.
Composition metadata caution | Directory paths point to another machine. Separate composition CSVs retain some '<0.005' qualifiers that the JSON represents as 0.005.
Recommended data matrix | (3200, 2048) intensities + sample-ID + measurement-file group ID
Train/test caution | Random row splitting can leak spectra from the same measurement file into both sets. Grouping by measurement file is more meaningful for cross-acquisition testing.
Physical-position caution | The files do not establish whether the five measurements are different physical positions or repeated acquisitions at one position.
Reference compositions for BAS/BCS samples are consolidated in the first workbook sheet (“Reference Materials”).

## Moving_Apr29
Sub-dataset – Medals_MovingSamples_2026-04-29
Moving-sample LIBS measurements acquired on 29 April 2026. The folder provides spectra and five JPEG images but no machine-readable acquisition-settings file.
Dataset overview
Property | Specification
Repository location | data/LSA/Medals_MovingSamples_2026-04-29
Sample IDs | 15
Measurement files per sample | 1
Total spectral files | 15 CSVs
Spectra per file | 5 or 6
Total spectra | 79
Wavelength channels per spectrum | 2048
Total intensity values | 161792
Supporting files | 5 JPEG images
Total directory size | Approximately 2.30 MiB
Acquisition date in filenames | 29 April 2026
Filename time span | 10:28:14–10:57:43; timezone unspecified
Organisation | All files directly in the folder; no sample subfolders
Moving-sample metadata | Folder name indicates moving samples, but motion speed, trajectory and pass mapping are not documented
Sample inventory
Sample ID | Material information | Spectra
469 | Ferritic stainless-steel reference ID | 5
470 | Ferritic stainless-steel reference ID | 5
471 | Ferritic stainless-steel reference ID | 5
472 | Ferritic stainless-steel reference ID | 6
473 | Ferritic stainless-steel reference ID | 5
551 | Phosphor-bronze reference ID | 6
552 | Phosphor-bronze reference ID | 5
553 | Phosphor-bronze reference ID | 6
554 | Phosphor-bronze reference ID | 5
555 | Phosphor-bronze reference ID | 5
556 | Phosphor-bronze reference ID | 6
M2-7_A3 | Sample label; composition not provided here | 5
SS405 | Sample label; composition not provided here | 5
SS465-1 | Sample label; composition not provided here | 5
SS467 | Sample label; composition not provided here | 5
By material-label family: 26 ferritic steel-reference spectra, 33 phosphor-bronze reference spectra, and 20 spectra from the four other sample IDs. All 15 IDs also occur in the Feb26 dataset; Al_ENAS-42000-1 is absent here.
CSV format and spectral axis
Property | Specification
Separator | Semicolon (;)
Header rows | 1
Data rows | 5 in 11 files; 6 in 4 files
Columns | 2050
MeasCtr | Sequential counter starting at 0, ending at 4 or 5
TimeStamp | Integer values; units and origin undocumented
Spectral intensity columns | 2048
Wavelength coverage | 341.952–896.916 nm
Wavelength ordering | Strictly increasing; identical across all files
Channel spacing | Nonuniform, approximately 0.242–0.294 nm
Mean channel spacing | Approximately 0.2711 nm
Intensity representation | Integers; physical units unspecified
Observed intensity range | 566–65,535
Optical spectral resolution | Not documented; should not be equated with channel spacing
Acquisition information and missing specifications
Item | Information
TimeStamp differences | 7,722–69,693 units; many intervals around 8,000–9,000 units. Units are not confirmed.
Settings file | None in this folder
Not established | Spectrometer model/serial number, exposure time, laser wavelength/energy/pulse duration/repetition rate, sample speed, motion geometry, number of passes, trigger synchronisation, measurement position, single-shot vs averaged rows
JPEG files | Five images are present; their filenames contain an April 15 date, which should not replace the April 29 acquisition date encoded in the CSV filenames.
Relationship to background-corrected data
Property | Specification
Matching corrected folder | data/LSA/Medals_April2026_BGCorrectedData
Filename relationship | Each of the 15 files has a matching '_removedBG.csv' counterpart
Additional corrected-file sample | The corrected folder additionally contains a 'Holz' file, absent from this moving-sample folder
Interpretation | Filename pairing supports original/corrected counterparts, but the correction algorithm is not documented in the folder metadata
Quality checks and analysis implications
Check / item | Result / implication
Consistent row widths | Every data row has 2,050 columns
Missing/non-numeric intensities | None found; all intensity entries parsed as integers
Exact duplicate files | None; all 15 CSV hashes differ
Values equal to 65,535 | 75 (approximately 0.046% of intensity values)
Spectra containing ≥1 value of 65,535 | 28 of 79 spectra (35.4%)
Samples without any 65,535 values | 471, 551, M2-7_A3
Saturation interpretation | Repeated maximum value suggests clipping; it may affect strong emission-line measurements
Recommended data matrix | (79, 2048) intensities + sample ID + measurement-file identifier
Split limitation | Only one acquisition file per sample ID; random row splitting tests within-acquisition performance rather than independent-acquisition generalisation
Cross-session potential | Overlapping labels and wavelength grid with Feb26 can support evaluation across acquisition sessions under moving-sample conditions, subject to confirmation of specimen identity/settings
Reference compositions for the BAS/BCS IDs are listed in the first workbook sheet. This folder itself contains no composition table.

## July2026
Sub-dataset – July2026_ScrapSamples
LIBS measurements acquired on 16 July 2026, containing scrap-sample identifiers and seven labelled European material grades.
Dataset overview
Property | Specification
Repository location | data/LSA/July2026_ScrapSamples
Sample groups | ScrapSamples and SteelGrades
Scrap sample files | 26 files containing 548 spectra
Industrial-grade files | 7 files containing 150 spectra
Total files | 33 CSVs
Spectra per file | 11–30, usually 20–24
Total spectra | 698
Wavelength channels per spectrum | 2048
Total intensity values | 1429504
Directory size | Approximately 6.90 MiB
Acquisition date in filenames | 16 July 2026
Filename time span | 13:50:42–17:29:29; timezone unspecified
Supporting metadata | No composition tables, settings files or photos in this folder
Organisation | One CSV per sample label
Scrap sample inventory
Sample | Spectra | Filename annotation / note
103-1 | 18
103-2 | 22
412-2 | 21
452 | 20
462 | 20
Band | 30
HF-1 | 20
HF-2 | 20
HF-4 | 20
HF-8 | 21
S2 | 23 | Einzelpulse_1ms: single pulses; includes '1ms' annotation
S3 | 20
S4 | 22 | Einzelpulse: single pulses
S5 | 22
S7 | 21
S8 | 21
S9 | 24
S10 | 23
S11 | 21
SR-1 | 22
SR-2 | 23
SR-11 | 21
SR-12 | 20
SR-13 | 23
SR-14 | 11 | Explicitly marked 'incomplete'
SS-345 | 19
Total | 548
The '1ms' annotation does not identify which acquisition parameter it describes and should not automatically be interpreted as exposure time.
Industrial-grade inventory
Filename label | Material number | Spectra
1-3974 | 1.3974 | 22
1-4301 | 1.4301 | 20
1-4404 | 1.4404 | 20
1-4462 | 1.4462 | 21
1-4539 | 1.4539 | 22
1-4860 | 1.4860 | 21
2-4605 | 2.4605 | 24
Total |  | 150
The folder contains material-number labels but no certificates or measured compositions. Published material-grade specifications are listed in the first workbook sheet.
CSV format and spectral axis
Property | Specification
Separator | Semicolon (;)
Header rows | 1
Data rows | 11–30 depending on file
Columns | 2050
MeasCtr | Starts at 0
TimeStamp | Integer values; units/reference origin not documented
Spectral intensity columns | 2048
Wavelength coverage | 341.952–896.916 nm
Wavelength ordering | Strictly increasing
Channel spacing | Approximately 0.242–0.294 nm, nonuniform
Mean channel spacing | Approximately 0.2711 nm
Intensity representation | Integers; physical units unspecified
Observed intensity range | 403–65,535
Wavelength-grid compatibility | Identical headers to Feb26 and Moving_Apr29 datasets
Optical spectral resolution | Not documented; cannot be inferred from channel spacing alone
Acquisition information and missing specifications
Item | Information
Information present | Filename dates/recording intervals and limited single-pulse annotations
Not documented | Spectrometer model/serial number; dataset-wide exposure time; laser model/wavelength/energy/current/repetition rate; trigger delay; collection geometry; sample movement; single-shot vs averaged rows; surface-position mapping; intensity calibration/background correction
TimeStamp increments | Irregular, 108–46,742 units; units and origin not documented
Cross-dataset caution | The February acquisition settings should not be assumed to apply to July
Quality checks and analysis implications
Check / item | Result / implication
Consistent row widths | All data rows have 2,050 columns
Missing/non-numeric intensities | None found; all parsed as integers
Exact duplicate CSVs within folder | None; all 33 hashes differ
Values equal to 65,535 | 3,520 values (0.246% of intensity values)
Spectra containing ≥1 value of 65,535 | 488 of 698 spectra (69.9%)
Explicitly incomplete sample | SR-14_incomplete, with 11 spectra
Duplicate datasets elsewhere | Byte-identical CSV collections exist in Industrial_Samples_2026-07 and Fw_ Scrap sample measurements
Saturation interpretation | Repeated maximum value strongly suggests clipping; although sparse by channel count, it affects roughly seven in ten spectra
Recommended data matrix | (698, 2048) intensities + sample-label + group columns
Train/test limitation | One file per sample label means file-group splitting cannot retain every label in both train and test. Random row splitting evaluates within-acquisition performance rather than independent specimens.

## Dataset Comparison
MEDALS Measured Sub-dataset Comparison
High-level comparison of the three measured LIBS sub-datasets documented in Deliverable D4.6.
Cross-dataset overview
Dataset | Acquisition date | Sample IDs / labels | CSV files | Total spectra | Channels | Wavelength range (nm) | Supporting metadata | Ground-truth availability | 65,535-affected spectra | Main use / interpretation | Main limitation
Medals_Samples_Feb26 | 11 Mar 2026 | 16 | 80 | 3200 | 2048 | 341.952–896.916 | Composition JSON, settings TXT, 3 JPEGs | BAS/BCS reference compositions for 469–473 and 551–556 | 492 / 3200 (15.4%) | Primary structured reference/acquisition dataset; repeated files per sample | Five files per sample, but physical-position relationship is not documented
Medals_MovingSamples_2026-04-29 | 29 Apr 2026 | 15 | 15 | 79 | 2048 | 341.952–896.916 | 5 JPEGs; no settings file | Reference identity can be linked to BAS/BCS IDs; no composition file in folder | 28 / 79 (35.4%) | Cross-session evaluation under moving-sample conditions | Only one file per sample; motion/acquisition settings not documented
July2026_ScrapSamples | 16 Jul 2026 | 33 file labels (26 scrap + 7 grades) | 33 | 698 | 2048 | 341.952–896.916 | No composition/settings/photos in folder | No sample-specific composition; 7 files carry European material-number labels | 488 / 698 (69.9%) | Industrial/scrap measurement collection with labelled grades | One file per label; limited acquisition metadata; strong saturation prevalence
Common spectral format
Property | Common characteristic
Wavelength channels | 2,048 in all three datasets
Wavelength coverage | 341.952–896.916 nm in all three datasets
Grid alignment | Identical wavelength headers across the three datasets
Channel spacing | Nonuniform, approximately 0.242–0.294 nm; mean ≈0.2711 nm
CSV metadata columns | MeasCtr and TimeStamp precede spectral intensity channels
Intensity representation | Integer values; no physical intensity unit documented
Spectral resolution caution | Channel spacing is not equivalent to optical spectral resolution
The three sub-datasets are directly compatible at the wavelength-column level, but differences in acquisition session, sample movement, settings metadata and saturation mean that raw intensity comparisons should be interpreted with care.